        - Appendix No. 2

        - ........................................ 								.............................................

        - Stamp of the Contractor									place and date

SUMMARY OF REQUIREMENTS

re. Delivery of ceilometer

CEILOMETER: ………………………………………………

Name of the device, producer, type/model

| Lp. | ## Subject | ## Description of requirements | Parameters offered |
| --- | --- | --- | --- |
| Technical characteristic of the device |  |  |  |
| 1 | Measuring principle | Optical (LIDAR) |  |
| 2 | Measuring range (CBH) | 5 m ... 15 000 m |  |
| 3 | Accuracy | ± 5 m |  |
| 4 | Range resolution | 5 m |  |
| 5 | Sampling rate | 100 MHz |  |
| 6 | NetCDF raw data resolution | 15 m (full range, compact file sizes) 5 m (5 m to 150 m range) |  |
| 7 | Time to measure | 2s … 600s (programmable) |  |
| 8 | Targets | Aerosols, clouds |  |
| 9 | Light source | Nd:YAG solid-state laser, wavelength 1064 nm |  |
| Interface and software characteristic |  |  |  |
| 1 | Standard interface | RS485, LAN |  |
| 2 | Optional interfaces | RS232 or Modem V.21, V.22, V.22bis |  |
| 3 | Communication | LAN Port: Web-Interface Serial Port |  |
| 4 | Software | For visualizing measured results |  |
| Electrical parameters |  |  |  |
| 1 | Power supply standard | 230 V(AC), ±10 % |  |
| 2 | Optional power supply | 110 VAC, 115 VAC, ±10 % |  |
| 3 | Power consumption | 250 W (standard) 800 W (in maximum heating mode) |  |
| 4 | UPS functionality | Internal backup battery for electronics, > 1 hrs |  |
| Operating conditions |  |  |  |
| 1 | Temperature | -40 °C ... +55 °C |  |
| 2 | Relative humidity | 0 % ... 100 % |  |
| 3 | Wind | 55 ms-1 |  |

Example of the device which meets the requirements above: Ceilometer CHM 15k „Nimbus” with accessories and software

REMARK

The Contractor should submit with the bid offer the Technical Specifications of offered devices (catalogs, brochures, catalog cards, etc.) confirming that the offered device meets the requirements described above.

....................................................................................

signature and stamp of the Contractor

or an authorized person